## Specification Comparison
|  | Temperature Control | Gas Control
Pecon Incubator (Zeis)
|  | Temp Module S | Heating Insert P Lab-Tek TM S1 | Incubator PM S1 |  | CO2 Module S | O2 Module S | Heating Device Humidity S
Range |  | Ambient to 60°C (Recommended 45°C) |  |  |  | 0.0 to 8.0 Vol-% | 0.0 to 21.0 Vol-% | 100% Relative Humidity
Accuracy |  | ± 0.15% of set point |  |  |  | ± 2% of set point | ± 0.05% of set point
Resolution |  | 0.01° |  |  |  | 0.01% vol | 0.01% vol
Dimensions (W x H x D mm) |  | 313 x 85.5 x 301 | 160 x 22 x 135 | 132 x 205 x 18 |  | 313 x 128 x 301 | 313 x 85.5 x 301 | 104 mm diameter, 124 mm height
Weight |  | ~5.4 kg |  |  |  | ~5.7 kg | ~5.7 kg | ~1.5kg
Method |  | 4 channel Direct to component and Gas Heating |  |  |  | Controlled release of CO2 Supply | Controlled release of N2 supply to displace O2 | Bubble the Gas flow through heated water to introduce humidity
Features |  | Supports heating of microscope objective | Supports Standard Slides as well as Petri dishes | Homogeneous Gas Distribution |  | Requires min 95% Concentrated pressurized CO2 supply @ 1-2 bar | Requires min 99% Concentrated pressurized N2 supply @ 1-2 bar | Accepts 250ml and 500ml humidification bottle
|  | Supports software to log parameters during operation | 4 Levelling screws | Non Hermetic Seal, Continuous displacement venting |  |  |  | Heated via Temp Module
|  |  | 46x21 mm observation opening | Heated via Temp Module
Heated via Temp Module
Supports perfusion experiments using 2 ports for tubes
General Working
Ibidi Silver Line Incubator
|  | Heating System Slide/Dish – Silver Line | Heated Lid (for Slide/Dish System) | Heated Plate (for Slide/Dish System) |  | Gas Incubation System, CO2 – Silver Line | Gas Incubation System, CO2 / O2 – Silver Line | Humidifying tower
Range |  |  | Ambient + 5°C to 50°C (minimum of 18°C) | Ambient + 5°C to 46°C (minimum of 18°C) |  | 0 to 15% vol | (CO2 specs the same) (O2: 0.5 to 21%) | 20 to 99% (lower bound dependent on environment and gas supply
Accuracy |  |  | ±0.2°C with a ±5°C Uniformity | ±0.2°C with a ±0.5°C Uniformity |  | ±0.2% vol + 3% of reading | (O2) ±0.2 % vol (typical) ±0.7% vol (max) (less accurate for 0.5% to 1% concentration) | Sensor Accuracy ±2.5%
Resolution | [5 empty] | 0.1% | 0.1% | 0.1%
Dimensions (W x H x D mm) |  | 230 x 90 x 170 | 127.5 x 20 x 85.5 | 127.5 x 15 x 85.5 |  | 230 x 90 x 170
Weight |  | 1.75kg |  |  |  | 2.23kg
Method |  | Output 6 individual heating channels |  |  |  | Controlled release of CO2 Supply into Pressurized Air Mixture | Controlled release of N2 supply to displace O2
Features |  | Compatible With Slide/Dish mount and with Multi well Mount | Levelling screws | Levelling screws |  | Gas outflow humidity can be controlled using Humidification Tower |  | Temperature control: 45°-55°
|  | Supports heating of microscope objective | Supports perfusion experiments using 2 ports for tubes |  |  | Humidity sensor attached to the Incubator for humidity control |  | Temperature Accuracy: ±2°C
|  | Supports software to log parameters during operation |  |  |  | Can heat the airflow tube to prevent condensation (Provides own Temperature control for Gas mixture) |  | Temperature Resolution: 0.1%
[6 empty] | Requires min 99% Concentrated pressurized CO2 supply And Compressed Air supply @ 0.8-1.2 bar | Requires min 99% Concentrated pressurized CO2 supply And min 99% concentrated pressurized N2 supply And Compressed Air supply @ 0.8-1.2 bar
Tokai STX Incubator
|  | STXG/STXF Temp and CO2 controller | Heated Lid For WSKM Incubator chamber designed for Zeis k-type frame | Heated Stage For WSKM Incubator chamber | Heated Water Bath For WSKM Incubator chamber | STXG/STXF Temp and CO2 controller | Optional separate STX-CO2O2 module | Humidity
Range |  |  | Ambient + 5°C to 65°C (minimum of 10°C) | Ambient + 5°C to 50°C (minimum of 10°C) | Ambient + 5°C to 50°C (minimum of 10°C) | 5% to 20% vol | (O2) 0.1% to 18%
Accuracy |  |  | ±0.3°C at heater surface (in specific conditions) | ±0.3°C at heater surface (in specific conditions) | ±0.3°C at heater surface (in specific conditions) | ±0.1% vol
Resolution |  |  | 0.1°C | 0.1°C | 0.1°C
Dimensions (W x H x D mm) |  | 160 x 256 x 280 |  |  |  |  | 160 x 256 x 270
Weight |  | 5.4kg |  |  |  |  | 4.9kg
Method |  | Combined Temperature and CO2 Control |  |  |  | Controlled release of CO2 supply | Controlled release of N2 supply to displace O2 Controlled release of CO2 | Forced humidity by heating reservoir of water in Incubation Chamber
Features |  | Supports direct sample temperature measurement | Uses Pt100 temperature sensor | Uses Pt100 temperature sensor | Uses Pt100 temperature sensor | Requires CO2 | Requires 100% CO2 and N2 supply at 1bar
|  | Supports add on CO2/O2 control unit for Oxygen control | Levelling screws | Levelling screws | Levelling screws |  | Overrides STXG/STXF Gas control
|  | STXG model utilises built in Gas mixer and 100% CO2 Supply to achieve CO2 control |  |  | Heated water bath provides forced humidity inside the chamber
STXF model utilises pre-mixed 5% CO2 Compressed air for Gas flow (No CO2 concentration control)
|  | Supports heating of microscope objective | WSKM Incubator chamber supports 3 different dish attachments For 35/50/60 ml dish, slides, well-plates
|  | Supports software to log parameters during operation | WSKM Incubator supports 3 access ports for sensor, 2 perfusion tubes And drug administration tube
Oko-Lab Incubator
|  | H301-T-UNIT-BL-PLUS | H301-K-FRAME |  |  | CO2-O2-UNIT-BL [0-20; 1-95] (CO2) | CO2-O2-UNIT-BL [0-20; 1-95] (O2) | Active Humidity Control
Range |  | Ambient +3° to 60° |  |  |  | 0 to 20% VOL | 0 to 95% VOL (using pure O2 supply, otherwise 0 to 20%) | • 85-95% @25 °C • 51-95% @37 °C • 26-95% @50 °C
Accuracy |  | 0.1° (for sample measurement mode) 0.3 for unmeasured sample |  |  |  | ±0.2% @ 5% | ±0.2% @ 5% | ±1%
Resolution |  | 0.1° |  |  |  | 0.1% vol | 0.1% vol | 1%
Dimensions (W x H x D mm)
Weight |  | 1.6kg |  |  |  | 3.5kg | 3.5kg
Method | [5 empty] | Controlled release of N2, CO2 and O2/Air | Controlled release of N2, CO2 and O2/Air
Features |  | Ambient temperature sensor for better control | Supports multiple lids for various experimental set ups |  |  | Can be connected with O2/CO2/N2 supply for full atmospheric contro | Can be connected with O2/CO2/N2 supply for full atmospheric contro
|  | Sample feedback mode allows for better sample temp control | supports multiple sample holder mounts for various culture vessels |  |  | Otherwise can be connected with CO2/Compressed air supply for partial oxygen control | Otherwise can be connected with CO2/Compressed air supply for partial oxygen control
|  | Allows for temperature cycling profiles | supports perfusion and microfluidic experiments

## Conceptual Design Specs
Potential Custom Solution
|  | Containment | Sterilization | Temperature | CO2 | O2 | Humidity | Optional
Range |  |  |  | Ambient to 55 | 0% to 15% | 0.5% to 20% | 90% to 99%
Accuracy |  |  |  | ±0.2°C | ±0.1% vol | ±0.1% vol |  | Required
Resolution |  |  |  | 0.1°C
Dimensions (W x H x D mm)
Weight
Method |  |  |  | Use electrical heating to heat the chamber
Features |  | Must be completely hermetically sealed | Must be disposable or Autoclave compatible
|  | Must support a Multi-well Plate (96well) | Must handle sterlisation solutions including: Bleach 70% ethenol
Should be operable in fume hood